| FG5 #237 ABSOLUTE GRAVITY MEASUREMENT |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Location: Hobart (Mt Pleasant VLBI observatory) |  |  |  |  |  |  |  |  |
| Date / time: (local/UTC) | 3/12/2013 |  | Observers: |  | Nick Dando |  |  |  |
| Station name: | Hobart |  |  |  |  |  |  |  |
| Station code/ID: | 9699.9160 |  | Polar motion offset x: |  |  | 0.0533” |  |  |
| Station mark number: | 9699.9160 |  | y: |  |  | 0.2962” |  |  |
| Latitude (N/S): |  |  | Source: |  | MGL (->Ser 7 predictions) |  |  |  |
| Longitude (E/W): |  |  | Setup height Upper: |  |  |  | 61.5 (mm) |  |
| MSL elevation (m): |  |  | Lower: |  |  |  | 59.0 (mm) |  |
| Gradient (uGal/cm): | ± |  | Total: |  |  |  | 12.05 (cm) |  |
| Transfer (datum) height: | 130 cm |  | Orientation: SS bubble |  |  |  | South |  |
| Laser power on @ / / h m |  |  | KRONOS: |  |  |  |  |  |
| Date / time: (local/UTC) | 3/12 14:30 AEDT | 4/12 11:20 AEDT |  |  | 5/12 9:15 |  |  | 6/12 8:15 |
| Spring Position (mV): | -16 | -120 |  |  | -350 > 6 |  |  | -153 |
| Ion pump (kV): | 3.82 | 3.85 |  |  | 3.85 |  |  | 3.84 |
| Ion pump (10-4A): | 1.25 | 0.15 |  |  | 0.04 |  |  | 0.01 |
| Fringe amplitude (mV): | 255 | 268 |  |  | 250 |  |  | 230 |
| Air temperature (°C): | 19.0 | 18.0 |  |  | 15.2 |  |  | 14.4 |
| Laser voltage: (V): | 3.02 | 2.9 |  |  | 2.89 |  |  | 2.79 |
| 1F voltages (V): D: | 1.82 | 1.84 |  |  | 2.00 |  |  | 2.04 |
| E: | 1.41 | 1.43 |  |  | 1.59 |  |  | 1.64 |
| F: | 1.05 | 1.07 |  |  | 1.20 |  |  | 1.27 |
| G: | 0.69 | 0.70 |  |  | 0.84 |  |  | 0.89 |
| Laser vertical check: | < 5 % | < 5 % |  |  | < 20 % |  |  | < 10 % |
| Level check: | DC: SS:  | DC: SS:  |  |  | DC: SS:  |  |  | DC: SS:  |
| GPS lock:  | # Sets ¦ drops/set ¦ set interval (m) ¦ drop interval (s) ¦ pulse delay (s) |  |  |  |  |  |  |  |
| ‘g’ soft. sync:  | / / / / |  |  |  |  |  |  |  |
| 240V power source: | Mains AC:  |  |  | Portable generator:  |  |  | Using UPS:  |  |
| Preliminary result: g = | 9 804 231 83.9 uGal |  |  | project file name: |  |  |  |  |
| Number of sets: | 66 sets |  |  |  |  |  |  |  |
| Set scatter: | 4.91 uGal |  |  | Station mark photo: |  |  |  |  |
| Measurement precision: | uGal |  |  | Setup photo: |  |  |  |  |
| Total uncertainty: | 1.93 uGal |  |  | Site photo: |  |  |  |  |
| Purpose of measurement: Vertical deformation |  |  |  |  |  |  |  |  |
| Data acquisition started @ / / h m |  |  |  | Data acquisition ended @ / / h m |  |  |  |  |
| Comments: 5/12 9:25 Re–verticalised no improvement in fringes. ~4AM 6/12 possible laser hop through laser temperature drifting, confirmed via re-processing and forcing “E” peak. |  |  |  |  |  |  |  |  |
| More comments: |  |  |  |  |  |  |  |  |